## Failure mode tests
Failure mode	Mitigation	Expected result	Testing
Analog inputs
ESD	Series resistor for current-limiting, capacitor for charge absorbing	Reading shift for unknown time, no damage to op-amps/ADC/others / X?? measured effect max. on control loop (?)	Apply ESD (???) while... / probing ADC input or getting continuous ADC readings, for bad-reading meas. / probing voltage across op-amp and ADC inputs, for protection-diode current meas.
EMI	Series resistor and capacitor (LPF)	No significant reading change, no damage to op-amps/ADC/others	Apply EMI (???) while... / probing ADC input or getting continuous ADC readings, for bad-reading meas. / probing voltage across op-amp and ADC inputs, for protection-diode current meas.
Sustained over-voltage (short to external)	Series resistor for current-limiting	Bad reading, but no damage to op-amps/ADC/others / Alarm from CPLD
Valve-induced noise	Good design, filtering	No significant reading change on valve switching	Switch all 4 valves at once while probing each ADC input

Valve drivers
CS pin internal short to +24V	10K series resistor to ADC, valve-movement-detect circuit	No damage to ADC or others; current-gain resistor may fail open / Alarm from CPLD	Add short from +24V to valve driver CS pin while probing ADC input
Driver fails open	Current sense output from valve driver, CPLD checks current after closing valve	Alarm from CPLD if current sense output is behaving correctly, stuck high, or stuck low / Failure to alarm if current sense is stuck at expected value	Disconnect valve
Driver fails closed	Current sense output from valve driver, CPLD checks current after opening valve	Alarm from CPLD if current sense output is behaving correctly, stuck high, or stuck low / Failure to alarm if current sense is stuck at expected value	Cut trace, short valve driver ctrl input high
Valves
Open coil	Current sense or fault output from valve driver	Alarm from CPLD	Disconnect one or multiple (?) valves
Shorted coil	Current sense or fault output from valve driver / Valve driver internal protection; PCB fuse for backup	No damage to valve driver, power supply, or others / No loss of system power / Alarm from CPLD	Add short across one or multiple (?) valves while… / probing valve driver current and +24V rail
Stuck fully/mostly open (expiration valve)	Valve-motion-detect circuit	Alarm from CPLD	??? while probing valve-motion-detect output
Stuck fully/mostly closed (expiration valve)	Valve-motion-detect circuit	Alarm from CPLD	??? while probing valve-motion-detect output
Stuck fully/mostly open (all other valves)	N/A (system-level)
Stuck fully/mostly closed (all other valves)	N/A (system-level)
I2C bus (all sensor readings)
ESD/EMI-induced noise	More of a software/firmware issue: / (comparatively) low bus speeds, software tolerance for occasional comms errors	Temporary bad reading but no damage	??? while probing both I2C lines, acquiring sensor readings (make sure noise happens during readout)
Valve-induced noise	Good design, filtering	No deviation from logic thresholds while valves switching	Switch all 4 valves at once while probing both I2C lines
Peripheral holds line low/high	More of a software/firmware issue: error checking, redundant buses	Alarm from CPLD or Pi, possible fail-over to redundant bus	Short one/both lines on either/both I2C bus low / Short one//both lines on either/both I2C bus high
Raspberry Pi
Software hang	More of a software/firmware issue: watchdog timer from CPLD
Incorrect state changes	More of a software/firmware issue: / valid state-transition checking, safety-combining with CPLD on valve outputs
Excessive current draw / fail short on 5V input	Separate +5VD supply, current-limited temperature-limited +5VD supply	+24V/5V may or may not temporarily go out of regulation, alarm from CPLD, no fire risk from 5V regulator / Possible power-ORing diode overheating under sustained overload	Short +5VD power rail while observing component temperatures, probing +24V and +5V, measuring +24V input current
Outputs fail open / power loss	Pull-down resistors on valve outputs		Disable +5VD regulator (disconnect enable input resistor to CPLD, short enable to gnd)
Outputs stuck high	Valid-state-transition checking from CPLD	Alarm from CPLD	Pi script to hold outputs high
Outputs stuck low	Safety-combining with CPLD on valve outputs	Valves remain closed	Pi script to hold outputs low
GPIO (inputs or outputs) fail shorted to 5V	Series resistors on CPLD/Raspberry Pi comms lines	No effect unless other failures (Pi behaves abnormally, etc.)
Power
Valve-induced sag	Good design, filtering	No significant sag in 24V power or downstream (+5V, +5VD)	Switch all 4 valves at once while probing +24V, +24VA, +5V, +5VD
24V input temporary surge	Bulk capacitance and TVS diode (TVS not used by default)	+24V stays within capacitor, 5V DC-DC input voltage ratings / +24VA stays within capacitor, sensor input voltage ratings / +5V, +5VD remain in regulation	??? while probing +24V, +24VA, +5V, +5VD
24V input sustained over-voltage	TVS diode and input fuse (neither populated by default)	Failure, possible fire risk / If populated: Fuse blows, but no fire risk (all components stay within ratings)	Apply over-voltage from 100W+ bench supply while probing +24V, +24VA, +5V, +5VD
24V temporary sag	Bulk capacitance on all rails	All voltages stay in regulation for ??? ms drop-out	Use relay to disconnect input for ??? ms while probing +24V, +24VA, +5V, +5VD
24V sustained under-voltage / (enough to run DC-DCs but potential valve problems)	Voltage sensing on CPLD ADC, / current sense output from valve driver (< expected on-state current), / valve-motion-detect circuit	Alarm from CPLD	Apply low continuous voltage(12V???) from bench supply
24V outage	Power-down alarm circuit on backup battery	Alarm from power-down alarm circuit	Disconnect 24V supply
5V supply failure (+5V)	Diode-ORed alarm/CPLD power from backup battery, / CPLD sensing of +5VA through ADC	Alarm from CPLD	Disable +5V regulator (short enable input to gnd)
5V supply failure (+5VD)	(Same as Raspberry Pi outputs fail open / power loss)
3.3V supply failure (+3.3V)	Power-down alarm circuit on backup battery	Alarm from power-down alarm circuit	Disable 3.3V regulator (cut trace and bring enable line low)
1.8V supply failure (+1.8V)	Power-down alarm circuit on backup battery	Alarm from power-down alarm circuit	Disable 1.8V regulator (cut trace and bring enable line low)
Short on +24V, +5V, +3.3V, +1.8V	Power-down alarm circuit on backup battery, current-limited power supplies, temperature-limited 5V supply	Power fails on multiple rails up to or including 24V input, / but alarm sounds from power-down alarm circuit, no fire risk / Possible power-ORing diode overheating under sustained overload	Short each power rail while observing component temperatures, probing +5V and +24V, measuring +24V input current
Short / excessive current draw on +5VD	(Same as Raspberry Pi excessive current draw / failt short on 5V)
Short / excessive current draw on +5VA	Resettable fuse, +5VA monitor (~+5VA_On signal)	Alarm from CPLD	Short +5VA while observing component temperatures, probing +5V and +24V, measuring +24V input current
Short / excessive current draw from +24VA / on a (0-10V) pressure sensor	(Same as Pressure sensor fails short on 24V power input)

NOTE: items in bold have unresolved questions, lines marked in green involve firmware/software issues, lines in gray are duplicates of issues in other categories

## Bring-up tests
Power rails	DC voltage	Noise (XXX Mhz BW)	Current draw (mA)	Conditions
+24V
+24VA
+5V
+5VA
+3.3V
+1.8V
+2.5V
+BATT
Power rail sequencing	ADC-supply waveform	+5VA waveform	Op-amp Vout waveform	ADC Vin waveform	Op-amp supply before ADC supply?	Max. ADC ESD-diode current?
+3.3V ADC vs. +5VA op-amp
+3.3V_RPi ADC vs. +5VA op-amp
Pressure sensor 1 input	Input	CPLD op-amp in	Pi op-amp in	CPLD ADC in	Pi ADC in	CPLD ADC reading	Pi ADC reading
Offset: short inputs, check voltages along signal chain and ADC reading
Gain: apply full-scale input, check voltages along signal chain and ADC reading
Pressure sensor 2 input
Offset: short inputs, check voltages along signal chain and ADC reading
Gain: apply full-scale input, check voltages along signal chain and ADC reading
O2 sensor input
Offset: short inputs, check voltages along signal chain and ADC reading
Gain: apply full-scale input, check voltages along signal chain and ADC reading
Valve driver basic functional: / disconnect Pi and short its outputs to +3.3V, selectively short +3.3V to CPLD valve outputs	LED on?	Output voltage waveform	CS pin waveform
1
2
3
4
Valve-motion circuit	CS pin waveform	Valve motion detect waveform
Normal turn-on: same as valve driver basic functional
Normal turn-off: same as valve driver basic functional
Power-down alarm	+3.3V threshold	Alarm tone freq.	Alarm repeat freq.	+BATT current waveform
Remove +3.3V power, wait for alarm to sound
Power-down detect circuits / Apply triangle wave to each rail in turn, check at which voltages the output changes state w/no glitching	Turn-off threshold	Turn-on threshold	Glitches?
+3.3V (power-down alarm)
+1.8V (power-down alarm)
+5VA (power-down alarm)
+5V (+BATT diode-OR circuit)
Audio amplifier / Apply square wave to each input in turn	Tone sounds?	+BATT current draw
Raspberry Pi input
CPLD input